## Slide 1: ST Calibration Update
Eric Pooser
03/18/2015

## Slide 2: Runs Used for the following Calibrations
FCAL Trigger
  Runs (1505 – 1525): ≈ 12 M events
  50 nA, 10 GeV, 1200 A, 10 mm CH2, 5.0 mm hole
  PC = Retracted, Radiator: 2 x 10-5 RL
FCAL/BCAL Trigger
  Runs (2397 – 2420): ≈ 125 M events
  100 nA, 10 GeV, 1200 A, 10 mm CH2, 5.0 mm hole
  PC = 5 x 10-3 RL, Radiator: 2 x 10-5 RL

## Slide 3: Time Walk Calibrations
Functional Form
  A ≡ Pedestal Corrected Pulse Integral
  A0 ≡ Mean of minimum ionizing Landau Peak
For now, A0 is hard coded to be 3000
  Nothing is gained by allowing this value to be a parameter
Future calibrations could implement a fitting procedure which would accurately feed A0 into the time walk fits on a channel by channel basis

## Slide 4: Pedestal Corrected Pulse Integral
Overflow
Dynamic Range of fADC: 0 – 30,000 channels

## Slide 5: Landau Peak Mean & Hit Times
Landau Peak (μ ≈ 3000 chan)

## Slide 6: Time Walk Corrections (Before & After)

## Slide 7: Time Walk Corrections (Before & After)

## Slide 8: Time Walk Corrections (Before & After)
σ ≈ 400 ps
σ ≈ 575 ps

## Slide 9: Time Walk Corrections (Before & After)

## Slide 10: Time Walk Corrections (CCDB)
Old TW Corrections
New TW Corrections

## Slide 11: Propagation Time Correction Event Selection
Only events where exactly 2 hits in the ST occurred were selected
  The two hits could not be in the same paddle
The events were then grouped into 5 possible scenarios such that the two hits occurred in the following sections of the ST:
  Straight section
  Bend section
  Nose section
  Bend & Nose section
  Straight, Bend, & Nose section

## Slide 12: Propagation Time Corrections (cont.)
The same data used for the TW corrections were used for the propagation time correction studies
Only FCAL/BCAL trigger runs using the new TW corrections are illustrated here
The time difference of the two hits are plotted against the difference of the hits location along the path of the scintillators

## Slide 13: Propagation Time Corrections (cont.)
The δt vs δz data were then fit with a polynomial of order 1 and is determined
The correction is then applied to the data
TW corrected and propagation time corrected time resolutions are obtained for selected regions of the Start Counter

## Slide 14: δt vs δz (Uncorrected)
Some of the data is contradictory to bench mark measurements made at FIU
This could simply be a lack of statistics

## Slide 15: δt vs δz (Corrected)
After the correction, the effective velocity of light in all regions of the ST become essentially infinite

## Slide 16: Time Resolutions for Runs (2397 – 2420)
The aforementioned propagation time corrections substantially improve the time resolution measurements
  The updated TW corrections provided for better time resolutions (≈ 40 ps )
These preliminary measurements indicate an average time resolution of the entire ST to be approximately 500 ps

## Slide 17: Future Plans
Further TW correction studies are needed
  Various functional forms need to be investigated
  Develop algorithm to locate mean of Landau peak in pedestal corrected pulse integral spectrum
Fall back on benchmark measurements made at FIU regarding propagation time and implement these into the PID library until we have more statistics from the upcoming Spring run
  A different algorithm is needed for the two cases (beam data vs. benchmark data)
Develop algorithm to measure propagation time corrections on a paddle by paddle basis
Develop algorithm to select events in 5 cm regions of the paddles so that we may investigate time resolution vs. distance plots so that we may compare beam data to FIU benchmark data