## Slide 1
Beamline Commissioning and Radiation Level
A. Somov, Jefferson Lab
E12-19-003
Experiment Readiness Review
May 7, 2020

## Slide 2: Outline
Beam Commissiong (charge 7)
Are the beam commissioning procedures and machine protection systems sufficiently defined for this stage?
Radiation level (charge 4)
What is the impact of the expected neutron radiation on GlueX detector components such as the SiPMs? Is any local shielding required? Are the radiation levels expected to be generated in the hall acceptable?

## Slide 3
Photon Beam Requirements
Experiment | Energy Range (GeV) | Polarization | Flux in the energy range of interest / sec | Flux on target (0.012 – 11.7 GeV) / sec
GlueX Design | 8.4 - 9.1 | 40 % | 108 | 1.8 109
GlueX II | 8.4 - 9.1 | 40 % | 5  107 | 9  108
This experiment | 8.4 - 9.1 | 40 % | 2  107 | 3.6  108
Electron beam current: 140 nA 350 nA
Radiator thickness: 4  10-4 X0 4  10-4 X0
Collimator diameter: 5 mm 5 mm
This experiment
GlueX II
Photon flux on target is about 2.5 smaller than GlueX II flux
( 5 times smaller than GlueX designed flux)

## Slide 4: Hall D Photon Beam Line
Use standard Hall D beam line equipment for SRC / CT !
~12 GeV
e-
75 m
Electron beam dump
photon beam
Photon
Beam dump
Radiator
GlueX detector
Tagger hodoscope
& microscope
Pair
Spectrometer
Collimators

## Slide 5: Targets
LD (LHe4) target. Standard GlueX target cell and cryogenic system
- LHe4 target was used in the Hall D PrimEx experiment
See C. Keith talk
Aluminum
(0.5 mm)
Copper
section (0.75 mm)
Carbon target
- Similar to PrimEx Be target

## Slide 6: SRC/CT Beamline Summary
12 GeV polarized photon beam has been successfully used by the GlueX
experiment in 2016 - 2019
Typical GlueX beam configurations:
- electron beam current 100 – 350 nA (data production)
< 5 nA (PS calibration)
Beamline equipment installed in Hall D is ready to use by the SRC/CT
experiment
Photon beam conditions used for the GlueX data production satisfied
specifications of the RC/CT experiment
Beam delivery/monitoring procedures and machine protection systems are
the same as for the GlueX experiment

## Slide 7
Radiation Level
Radiation damage to Silicon Photomultipliers (SiPM)
produced by neutron background
Radiation level in Hall D
Charge 4

## Slide 8
SRC Targets and Run Period
Target | Thickness |  |  | Atoms/cm2 | Run Time (days)
 | cm | g/cm2 | %X0 |  | 
LH | 30 | 2.1 | 3.4 | 1.281024 | GlueX
D | 30 | 4.9 | 4.1 | 1.511024 | 4.5
4He | 30 | 3.5 | 4.0 | 5.681023 | 1
12C | 1.5 | 3 | 7 | 1.451023 | 7
SRC/CT photon flux is 5 times smaller than GlueX designed flux
The largest neutron background will be produced on liquid deuterium and
helium targets

## Slide 9
Radiation Damage to SiPM
- displacement of atoms in the lattice
associated with the kinetic energy
release to matter
- characterized by damage function
(similar to biological damage curve)
- in simulation, convert particle fluence
to the equivalent fluence of 1 MeV
neutrons using the damage function
- largest damage is produced by neutrons
with Ekin ≥ 1 MeV
Increase of the SiPM dark current
1 MeV

## Slide 10
Silicon Photomutipliers in Hall D
array of 3 mm x 3 mm SiPMs (3840 sensors)
calorimeter design allows for a factor of 5 increase of the dark rate
- dark current decreases by about a factor of 3 when SiPMs
are cooled to 5 C
radiation hardness studied using Am-Be neutron source and
beam test setup in Hall A
- https://halldweb.jlab.org/wiki/index.php/SiPM_Radiation_Hardness_Tests
- Radiation hardness test of BCAL SiPMs, NIM A 698 (2013) 234
- Radiation tolerance survey, Proceedings of IEEE2012 conference
- dark current is doubled after irradiating SiPMs with the equivalent fluence of about 108 nEQ / cm2
(critical dose about 32 rem, dark current increase by x5 )
expected lifetime at GlueX designed luminosity (108  / sec)
5 years for Hydrogen
3.2 years for Helium
(assuming the experiment operation efficiency of 33%)
Barrel Calorimeter

## Slide 11
Silicon Photomutipliers in Hall D
Start Counter
3 mm x 3 mm SiPMs (120 sensors)
each ST paddle is instrumented with 4 sensors
Measure similar dependence of the dark rate on the neutron radiation
as for the BCAL sensor
scintillator detector, not very sensitive to the dark current. Large
readout thresholds can be set
- signal pulse amplitude > 100 pixels
- single pixel dark rate: 900 kHz per SiPM at room temperature
(number of noise pixels in 100 ns time window - 0.4, pulse width – 40 ns)
Increase of the dark rate by a factor of 10 is not a problem

## Slide 12
Barrel Calorimeter SiPMs ( 4 x 4 array of 3 mm x 3 mm)
Start Counter SiPMs ( 3 mm x 3 mm )
BCAL SiPMs
ST SiPMs
Silicon Photomutipliers in Hall D
BCAL SiPMs

## Slide 13
Simulation of the Radiation Level in Hall D
Estimated background from liquid hydrogen, helium, and deuterium targets using
Fluka and Geant simulations provided by the JLab Radiation Control group

## Slide 14
Neutron Background Induced by LH and LHe Targets
SRC/CT beam flux on target is 5 times smaller than that for GlueX
Background induced by the LHe target will not exceed the GlueX level
Background was evaluated for the GlueX
designed luminosity
In the BCAL, background is larger In the
downstream end
- it increases by about 30 % for the
LHe target
Background is significantly larger in the
Start Counter SiPM region

## Slide 15
He
D
Recent calculations by Pavel Degtiarenko
Computed using RadCon Geant and Fluka
- relatively good agreement between Geant and Fluka
Neutron Flux: Liquid He and D Targets
Neutron flux originated from LHe and LD targets

## Slide 16
Neutron Flux: Liquid He and D Targets
Deuterium
Helium
ST
BCAL downstream
BCAL upstream
Neutron angular distributions, FLUKA simulation

## Slide 17
1 MeV Equivalent Flux: He and D Targets
1 MeV neutron equivalent fluence for the Deuterium target is
larger than that for the He target by about a factor of
- 3 – 4 for the Start Counter and BCAL downstream regions
- 2 for the BCAL upstream SiPM region
D
He

## Slide 18
Summary: Radiation Damage to SiPM
 | GlueX Design | SRC/CT | 
 | LH target | LHe4 target | LD target
BCAL | 5 years | 19.9 years | 6.5 years
ST | 2.1 years | 2.6 years | 245 days
Expected life time of SiPMs
Estimate life time using radiation doses computed by RadCon
Conservative assumptions regarding neutron background induced by deuterium target
Assume experiment operation efficiency of 33 %
Degradation of the ST SiPMs after the run with deuterium target is equivalent to taking data in the GlueX experiment for about 1.4 months
SRC/CT run time 1x3 = 3 days 4.5x3 = 13.5 days

## Slide 19
Monitoring SiPM Dark Current
We are planning to monitor the dark rate of the Start Counter SiPMs
during the SRC/CT experiment
- special 10 min runs with the random trigger and unsparcified readout
of fadc250
- measure pedestal broadening due to the increase of the dark rate
Checked pedestals for ST during PrimEx run (20 PAC days on He4 target,
2.8 times smaller photon flux than in our experiment)
- no significant increase of the pedestal width observed after taking
data ( < 10 %, on the level of systematic uncertainties)

## Slide 20
Radiation Level in Hall D
Estimated neutron dose equivalent rate in Hall D(ceiling) produced in the deuterium
target is expected to be 3 – 4 times larger than that for helium (conservative assumption)
He
D
Dose Equivalent Rate: He and D Targets

## Slide 21
Radiation Level in Hall D
Experimental Hall D was designed to handle the photon flux about 2 times larger
than the GlueX nominal flux (corresponding to 108  /sec in the energy range
between 8.4 GeV and 9.1 GeV )
The SRC/CT experiment will run at 5 times smaller photon flux
The largest neutron background is expected to be produced on the deuterium
target. This background was estimated to be on the level of the designed GlueX
Background in the experimental Hall will be monitored by RadCon

## Slide 22
Monitoring Radiation Level in Hall D
Installed quick access ionization chambers (show locations)
Install thermoluminescent dosimeters (TLD) close to the target
Install Bonner spheres (determine the energy spectrum of neutons) close
to the target (coordinate with RadCon group)

## Slide 23
Radiation Level Summary
The SRC/CT experiment will be operated at the photon flux 5 times smaller than that of the designed GlueX experiment
Neutron background was estimated using Fluka and Geant simulations provided by the RadCon group
Degradation of the SiPM performance due to the neutron radiation, the dark rate increase, is expected to be small even for the deuterium target (for the ST, the degradation is equivalent to running GlueX for 1.4 months).
SiPM dark current will be monitored during the SRC/CT experiment
No local shielding is required for the SRC/CT experiment
The radiation level in Hall D is expected not to exceed the GlueX designed level
The background level in Hall D will be continuously monitored by the RadCon group

## Slide 24: Backup Slides

## Slide 25: Photon Flux Measurements with Pair Spectrometer
Reconstruct the energy of a beam photon by detecting e pairs
- measure the photon beam flux and spectrum of the collimated photon beam
Two layers of scintillator detectors:
High-granularity hodoscope
(measure photon energy in the
range 6 – 12 GeV )
Low-granularity counters (use in trigger)
Amorphous Radiator
Polarized Beam

## Slide 26: Tagged Photon Beam
Microscope
array 100 x 5 of thin scintillators
detect electrons in the energy
range 8.1 GeV < E  < 9.2 GeV
Fixed Array Hodoscope
Photons to Hall D
Electrons to the
beam dump
Radiator
12 GeV
e beam
cover large energy range 3 GeV < E  < 11.7 GeV,  E / E  < 0.002
233 counters installed
detect tagged ellectrons with E > 9.2 GeV during data runs
accidental background for PrimEx-D < 4 %

## Slide 27
Comparison between FLUKA and RadCon Geant
He
He
D
D
Geant
FLUKA
P. Degtyarenko
P. Degtiarenko

## Slide 28: Photon Flux Measurements with Pair Spectrometer
PS acceptance and energy determination
- Ray tracing (measure magnetic field map)
- Calibrate using total-absorption-counters (TAC) at low luminosity
Special trigger type – continuous flux monitor ( typical rate 1 – 3 kHz)
- fadc / CTP scalers inserted to the data stream and EPICS
3.4 mm collimator ,
2 10-5 X0 radiator,
2 nA beam current
Calibration is performed regularly; takes 2-6 hours
- will be performed during PrimEx-D run
Monitor the photon flux with the precision < 1 %

## Slide 29: Target Density Monitor
Short term stability control:
- photon beam flux provided by the PS
- rates in the Start Counter (ST) and Time-of-Fight (TOF) wall
Long term stability control:
- monitor using Compton process; expected rate in the photon range of
interest is about 30 Hz
ST rate dependence on the target density
ST consists of 30 paddles surrounding the target
ST rate for production runs: 250 kHz / paddle
Coincidence of hits between the ST and TOF
(2 x 2 bars in TOF at R = 30 cm & one ST paddle)
1.5 kHz

## Slide 30
Pair Spectrometer Acceptance Calibration
Calibrate PS acceptance using total-absorption counter (TAC)
Data samples were acquired for three converters
- 5.7 x 10-3 R.L. Al (508 m foil)
- 2.1 x 10-3 R.L. Be (750 m foil)
- 0.21 x 10-3 R.L. Be (75 m foil)
Run conditions
- 3.4 mm collimator , 2 10-5 radiator, 2 nA beam current
Trigger
- run two triggers in parallel: PS and TAC (energy sum)
- PS rate: 10 Hz 750 m Be
- TAC rate: 200 – 300 kHz ( trigger prescaling factor 129 )

## Slide 31
Summary: Radiation Damage to SiPM
 | GlueX Design | SRC/CT | 
 | LH target | LHe4 target | LD target
BCAL | 8 years | 31 years | 10.1 years
ST | 3.1 years | 3.8 years | 1 years
Expected life time of SiPMs
Estimate life time using radiation doses computed by RadCon
Conservative assumptions regarding neutron background induced by deuterium target
Assume experiment operation efficiency of 33 %
SRC/CT run time 1x3 = 3 days 4.5x3 = 13.5 days
Estimates based on the damage described in the NIM paper
(more conservative estimates are in the main talk)